## Training Data
Tag | Patterns | Response / Troubleshooting Steps
vibro_not_working_no_led | VIBRO is not working and LED is not on Vibro power issue, no LED My vibro won't turn on | To fix the VIBRO (NOT ON LED), follow these steps: 1. CHECK THE FUSE -> REPLACE FUSE. 2. CHECK THE POWER SUPPLY -> REPLACE POWER SUPPLY. 3. CHECK THE SUPPLY VOLTAGE (11v TO 14V) -> REPLACE POWER SUPPLY. 4. CHECK THE GUN ELEMENT SHORT -> REPLACE THE GUN. 5. REPLACE THE DC CONNECTOR. 6. CHECK THE STIRRER BOARD -> REPLACE THE STIRRER BOARD.
vibro_led_on_off | VIBRO is not working but LED just goes on and off Vibro LED blinking but not working Vibro LED flashes on and off | To troubleshoot VIBRO (LED JUST ON AND OFF): 1. CHECK THE GUN ELEMENT SHORT -> REPLACE THE GUN. 2. CHECK DC CONNECTOR -> REPLACE THE DC CONNECTOR. 3. CHECK THE STIRRER BOARD -> REPLACE THE STIRRER BOARD.
vibro_no_vibration | Machine is on but vibration is not present Vibro on but no vibration Vibration stopped working on vibro | To resolve VIBRO (MACHINE ON BUT NO VIBRATION): 1. CHECK THE SUPPLY VOLTAGE (11v TO 14V) -> REPLACE POWER SUPPLY. 2. CHECK THE GUN ELEMENT SHORT -> REPLACE THE GUN. 3. CHECK THE FREQ.POT AND CABLE -> REPLACE THE FREQ.POT. 4. CHECK THE STIRRER BOARD -> REPLACE THE STIRRER BOARD.
vibro_continuous_vibration | Machine continuously vibrating Vibro won't stop vibrating Continuous vibration issue | To fix VIBRO (MACHINE CONTINIOUSLY VIBRATING): 1. CHECK THE TIMER POT. AND CABLE -> REPLACE THE TIMER POT. 2. CHECK THE STIRRER BOARD -> REPLACE THE STIRRER BOARD.
vibro_low_vibration | Low vibration on the machine Vibro vibration is very weak Vibration is not strong enough | To fix VIBRO (LOW VIBRATION): 1. CHECK THE SUPPLY VOLTAGE (11v TO 14V) -> REPLACE POWER SUPPLY. 2. CHECK THE FREQ.POT AND CABLE -> REPLACE THE FREQ.POT. 3. CHECK THE STIRRER BOARD FREQ.TRIMPOT ADJUST -> ADJUST THE FREQ.TRIMPOT. 4. CHECK THE STIRRER BOARD -> REPLACE THE STIRRER BOARD.
vibro_vibration_led_not_working | Vibration indication LED is not working Vibro indication light is off Vibration LED is dead | To troubleshoot VIBRO (VIBRATION INDICATION LED NOT WORKING): 1. CHECK THE LED IS WORKING AND CHECK THE CONNECTING CABLE QUALITY -> REPLACE THE LED AND CABLE ASSY. 2. CHECK THE VOLTAGE OF THE INDICATION LINE -> REPLACE THE STIRRER BOARD.
vibro_power_led_not_working | Power LED is not working on vibro Main power light is out Vibro power LED off | To troubleshoot VIBRO (POWER LED NOT WORKING): 1. CHECK THE LED IS WORKING AND CHECK THE CONNECTING CABLE QUALITY -> REPLACE THE LED AND CABLE ASSY. 2. CHECK THE VOLTAGE OF THE INDICATION LINE -> REPLACE THE STIRRER BOARD.
vibro_stirrer_only_on_adapter_plugging | Stirrer working only during adapter plugging time Vibro not working by the push switch Push switch not starting vibro stirrer | To fix VIBRO (STIRRER WORKING ONLY THE ADAPTER PLUGGING TIME / PUSH SWITCH FAIL): 1. CHECK THE KEY FUNCTIONALITY AND CABLE ASSY. QUALITY -> REPLACE THE KEY. 2. CHECK THE STIRRER BOARD -> REPLACE THE STIRRER BOARD.
solar_charger_no_charging_voltage | Not present the battery charging voltage - 0V Solar charger board voltage is zero Battery is not getting charged, 0V | To fix SOLAR CHARGER BOARD (BATTERY CHARGING VOLTAGE - 0V): 1. CHECK THE CHARGER ADAPTER VOLTAGES - 12.5V AND 24V -> REPLACED THE CHARGER ADAPTER. 2. CHECK THE CHARGER FUSE -> REPLACE THE FUSE. 3. CHECK THE CHARGER ADAPTER TO SOLARCHARGER BOARD WIRING. 4. CHECK THE SOLAR CHARGER BOARD COMPONENTS DAMAGE -> REPLACE THE SOLAR CHARGER BOARD.
solar_charger_low_voltage | Battery charging voltage is low Solar charger output is low | To fix SOLAR CHARGER BOARD (BATTERY CHARGING VOLTAGE IS LOW): 1. CHECK THE CHARGER ADAPTER VOLTAGES - 12.5V AND 24V -> REPLACED THE CHARGER ADAPTER. 2. REMOVE THE BATTERY AND CHECK THE VOLTAGE ADJUST TRIM.POT. OF THE SOLAR CHARGER BOARD -> REPLACE THE CHARGER BOARD.
solar_charger_high_voltage | Battery charging voltage is too high Solar charger overcharging | To fix SOLAR CHARGER BOARD (BATTERY CHARGING VOLTAGE IS TOO HIGH): 1. REMOVE THE BATTERY AND CHECK THE VOLTAGE ADJUST TRIM.POT. OF THE SOLAR CHARGER BOARD -> REPLACE THE CHARGER BOARD.
solar_charger_battery_not_charging | Battery charging voltage is correct but battery not charging Voltage is fine but battery won't charge | To fix SOLAR CHARGER BOARD (VOLTAGE CORRECT BUT BATTERY NOT CHARGING): 1. CHECK THE BATTERY CONDITION -> REPLACE THE BATTERY. 2. CHECK THE SOLAR CHARGER BOARD COMPONENTS DAMAGE -> REPLACE THE SOLAR CHARGER BOARD.
solar_charger_led_not_working | Battery charging indication LED not working Solar charger LED broken | To fix SOLAR CHARGER BOARD (BATTERY CHARGING INDICATION LED NOT WORKING): 1. CHECK THE BATTERY CONDITION -> REPLACE THE BATTERY. 2. CHECK THE LED IS WORKING AND CHECK THE CONNECTING CABLE QUALITY -> REPLACE THE LED AND CABLE ASSY. 3. CHECK THE VOLTAGE OF THE CURRENT INDICATION LINE -> REPLACE THE SOLAR CHARGER BOARD.
solar_charger_low_voltage_cutoff_not_working | Low voltage cutoff not working Cutoff relay failing | To fix SOLAR CHARGER BOARD (LOW VOLTAGE CUTOFF NOT WORKING): 1. CHECK THE LOW VOLTAGE CUTOFF RELAY(K2) IS WORKING(POWER ON SWITCH SUDDEN ON/OFF/ON) -> REPLACE THE SOLARCHARGER BOARD.
solar_charger_cutoff_led_not_working | Low voltage cutoff LED not working | To fix SOLAR CHARGER BOARD (LOW VOLTAGE CUTOFF LED NOT WORKING): 1. CHECK THE LED IS WORKING AND CHECK THE CONNECTING CABLE QUALITY -> REPLACE THE LED AND CABLE ASSY. 2. CHECK THE VOLTAGE OF THE CUTOFF INDICATION LINE -> REPLACE THE SOLAR CHARGER BOARD.
solar_charger_machine_full_time_battery | Machine full time working from battery but battery not full Always on battery power | To fix SOLAR CHARGER BOARD (MACHINE FULL TIME WORKING FROM BATTERY): 1. CHECK THE CHARGER ADAPTER VOLTAGES - 12.5V AND 24V -> REPLACE THE CHARGER ADAPTER. 2. CHECK THE SOLAR CHARGER BOARD LOW VOLTAGE SETTING TRIM POT. -> REPLACE THE SOLARCHARGER BOARD. 3. CHECK THE CHARGER ADAPTER DETECTION RELAY(K1) IS WORKING -> REPLACE THE SOLARCHARGER BOARD.
solar_charger_not_working_from_external_battery | Machine not working from the external connected battery | To fix SOLAR CHARGER BOARD (NOT WORKING FROM EXTERNAL BATTERY): 1. CHECK THE EXTERNAL BATTERY VOLTAGE ABOVE 11.5V -> CHARGE THE EXTERNAL BATTERY. 2. CHECK THE RELAY(K1) IS WORKING (USE THE CHARGER ADAPTER) -> REPLACE THE SOLAR CHARGER BOARD. 3. CHECK THE LOW VOLTAGE CUTOFF TRIM POT. OF THE SOLAR CHARGER BOARD -> REPLACE THE SOLAR CHARGER BOARD.
compact_adapter_output_zero | Adapter output voltage is zero LSE V3 ECO V Compact adapter has 0V output | To fix COMPACT ADAPTER (OUTPUT VOLTAGE IS ZERO): 1. CHECK THE STATUS LED IS GLOWING. 2. OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT). 3. CHECK THE AC CORD -> REPLACE AC CORD. 4. CHECK THE ADAPTER PCB -> REPLACE THE ADAPTER PCB.
compact_adapter_output_zero_led_on | Adapter output voltage is zero but LED is glowing 0V output but adapter light is on | To fix COMPACT ADAPTER (OUTPUT VOLTAGE ZERO BUT LED GLOWING): 1. CHECK THE OUTPUT DC CONNECTER OF THE ADAPTER -> REPLACE THE OUTPUT DC CABLE.
compact_adapter_low_voltage | Adapter output voltage showing 8v to 9v Compact adapter low voltage output | To fix COMPACT ADAPTER (OUTPUT SHOWING 8V TO 9V): 1. OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT) -> REPLACE THE ADAPTER PCB.
charger_adapter_output_zero | Charger adapter output voltage is zero LSE-S V3 ECOD ECOSV No voltage from charger adapter | To fix CHARGER ADAPTER (OUTPUT VOLTAGE ZERO): 1. CHECK THE STATUS LED IS GLOWING. 2. OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT). 3. CHECK THE AC CORD -> REPLACE AC CORD. 4. CHECK THE CHARGER ADAPTER PCB -> REPLACE THE CHARGER ADAPTER PCB.
charger_adapter_output_zero_led_on | Charger adapter output voltage is zero but LED is glowing | To fix CHARGER ADAPTER (OUTPUT ZERO BUT LED GLOWING): 1. CHECK THE OUTPUT DC CONNECTER OF THE ADAPTER -> REPLACE THE OUTPUT DC CABLE.
charger_adapter_12v_low | Adapter output 12 voltage showing 8V to 9V but 24V is 24V 12V line is low on charger adapter | To fix CHARGER ADAPTER (12V SHOWING 8V TO 9V AND 24V IS 24V): 1. OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT) -> REPLACE THE CHARGER ADAPTER PCB.
charger_adapter_24v_low | Adapter output 12 voltage showing 12V and 24V is low voltage or zero | To fix CHARGER ADAPTER (12V OK, 24V LOW OR ZERO): 1. OPEN THE ADAPTER AND CHECK PCB QUALITY( COMPONENTS BURN OR NOT) -> REPLACE THE CHARGER ADAPTER PCB.
ecod_dpst_please_wait | System on time please wait notification continuously showing Stuck on please wait screen | To fix ECOD-DPST BOARD (PLEASE WAIT CONTINUOUSLY SHOWING): 1. CHECK THE SD CARD IS OK -> REPLACE THE SD CARD.
ecod_dpst_display_blank | Display is blank on DPST board Blank LCD screen | To fix ECOD-DPST BOARD (DISPLAY BLANK): 1. CHECK THE DPST BOARD BREATHING LED IS BLINKING -> REPLACE THE DPST BOARD. 2. CHECK THE DIRECTION OF THE DISPLAY FRC CABLE CONNECTOR -> CONNECT THE FRC CONNECTOR AT OPPOSITE DIRECTION. 3. CHECK THE DISPLAY FRC CABLE QUALITY AND LCD QUALITY -> REPLACE THE DISPLAY.
ecod_dpst_display_half | Display shown half portion only Only half the screen is working | To fix ECOD-DPST BOARD (DISPLAY SHOWN HALF PORTION ONLY): 1. CHECK THE DPST TO LCD FRC CABLE AND CONNECTOR QUALITY -> REPLACE THE DPST TO LCD CABLE. 2. CHECK THE DISPLAY MODULE FUNCTIONALITY -> REPLACE THE LCD MODULE. 3. CHECK THE DPST BOARD FUNCTIONALITY -> REPLACE THE DPST BOARD.
ecod_dpst_rtc_reset | RTC time not saved Restart time time changed or re-setted | To fix ECOD-DPST BOARD (RTC TIME NOT SAVED): 1. CHECK THE CMOS BATTERY VOLTAGE ABOVE 2V -> REPLACE THE CMOS BATTERY. 2. CHECK THE U5 AND PROCESSOR CONNECTING LINE SOLDERING -> RE-SOLDER THAT CONNECTION.
ecod_dpst_rtc_not_running | RTC time not running Clock is frozen | To fix ECOD-DPST BOARD (RTC TIME NOT RUNNING): 1. CHECK THE U5 IC CRYSTAL - Y3 AND ALSO CHECK C37 & C38 -> REPLACE THE Y3 CRYSTAL AND C37&C38.
ecod_dpst_keypad_not_working | DPST key pad not working Buttons not responding on DPST | To fix ECOD-DPST BOARD (DPST KEY PAD NOT WORKING): 1. CHECK THE DIRECTION OF THE KEYPAD FRC-10 CABLE CONNECTOR -> CONNECT THE FRC-10 CONNECTOR AT OPPOSITE DIRECTION. 2. CHECK THE DPST KEYPAD QUALITY -> REPLACE THE DPST KEYPAD. 3. CHECK THE DPST BOARD QUALITY -> REPLACE THE DPST BOARD.
ecod_dpst_usb_error | USB initialization error shown USB error on screen | To fix ECOD-DPST BOARD (USB INITIALIZATION ERROR): 1. CHECK THE DPST TO BACKPANEL CABLE QUALITY AND DIRECTION OF THE FRC CONNECTOR -> REPLACE THE FRC CABLE OR CONNECT OPPOSITE. 2. CONNECT THE NEW BACK PANEL BOARD AND CHECK -> REPLACE THE DPST BOARD.
ecod_dpst_weighing_scale | Weighing scale not working Scale disconnected | To fix ECOD-DPST BOARD (WEIGHING SCALE NOT WORKING): 1. CHECK THE REQUIRED SETTINGS IS APPLIED -> SET THE REQUIRED SETTINGS VALUES. 2. CHECK THE SCALE INTERFACING CABLE CONNECTIVITY -> REPLACE THE CABLE. 3. CHECK THE BACK PANEL TO DPST SCALE CABLE -> REPLACE THE CABLE. 4. CHECK THE BACKPANEL PCB FUNCTANALITY -> REPLACE PCB. 5. CHECK THE DPST BOARD FUNCTANALITY -> REPLACE THE DPST BOARD.
ecod_dpst_computer_output | Computer output not present | To fix ECOD-DPST BOARD (COMPUTER OUTPUT NOT PRESENT): 1. CHECK THE DATAOUTPUT FORMAT OF THE ANALYZER SETTINGS -> SET REQUIRED VALUES. 2. CHECK THE MACHINE COMPUTER OUT TO COMPUTER CABLE CONNECTIVITY -> REPLACE CABLE. 3. CHECK THE ANALYZER MAIN BOARD TO BACKPANEL INTERFACING CABLE -> REPLACE CABLE. 4. CHECK THE BACKPANEL PCB FUNCTANALITY -> REPLACE PCB. 5. CHECK THE ANALYZER MAIN BOARD FUNCTANALITY -> REPLACE MAIN BOARD.
ecod_dpst_external_display | External display data not present | To fix ECOD-DPST BOARD (EXTERNAL DISPLAY DATA NOT PRESENT): 1. CHECK THE MACHINE EXTERNAL DISPLAY OUT TO DISPLAY CABLE CONNECTIVITY -> REPLACE CABLE. 2. CHECK THE DPST BOARD TO BACKPANEL EXTERNAL DISPLAY CABLE -> REPLACE CABLE. 3. CHECK THE BACKPANEL PCB FUNCTANALITY -> REPLACE PCB. 4. CHECK THE DPST BOARD FUNCTANALITY -> REPLACE DPST BOARD.
ecod_dpst_rate_chart | Rate not taken from the chart Rate not viewable | To fix ECOD-DPST BOARD (RATE NOT TAKEN FROM THE CHART / NOT VIEW): 1. CHECK THE SD CARD IS OK -> REPLACE THE SD CARD. 2. CHECK THE RATE CHART FILE IS SAVED IN THE SD CARD. 3. CHECK THE SAVED FILE TYPE AND DATAS FORMAT IN CSV -> SAVE TO THE CORRECT FORMAT AND TYPE (CSV). 4. CHECK THE DPST BOARD FUNCTANALITY -> REPLACE THE DPST BOARD.
ecod_dpst_farmer_details | Farmer details not shown | To fix ECOD-DPST BOARD (FARMER DETAILS NOT SHOWN): 1. CHECK THE SD CARD IS OK -> REPLACE THE SD CARD. 2. CHECK THE RATE CHART FILE IS SAVED IN THE SD CARD. 3. CHECK THE SAVED FILE TYPE AND DATAS FORMAT IN CSV -> SAVE AS CSV. 4. CHECK THE DPST BOARD FUNCTANALITY -> REPLACE THE DPST BOARD.
ecod_dpst_analyzer_connectivity | Analyzer connectivity error shown | To fix ECOD-DPST BOARD (ANALYZER CONNECTIVITY ERROR): 1. CHECK THE ANALYZER TO DPST BOARD INTERFACING CABLE AND CONNECTOR QUALITY -> REPLACE CABLE. 2. CHECK THE ANALYZER MAIN BOARD FUNCTANALITY -> REPLACE MAIN BOARD. 3. CHECK THE ANALYZER DPST BOARD FUNCTANALITY -> REPLACE DPST BOARD.
ecod_dpst_wifi_gsm_error | WiFi GSM module error shown Cellular or wifi not connecting | To fix ECOD-DPST BOARD (WIFI/GSM MODULE ERROR SHOWN): 1. CHECK THE WIFI/GSM SETTINGS DETAILS -> ENTER THE REQUIRED SETTINGS. 2. CHECK THE DPST TO BACKPANEL FRC CABLE AND CONNECTION DIRECTION -> REPLACE OR RECONNECT FRC CABLE. 3. CHECK THE WIFI/GSM MODULE BREATHING LED IS WORKING -> REPLACE THE WIFI/GSM MODULE. 4. CHECK THE ANTENNA SMA CABLE CONNECTION AND CONNECTOR -> REPLACE THE ANTENNA AND SMA CABLE.
ecod_dpst_sms_not_sending | SMS not send to the farmer Messages failing to send | To fix ECOD-DPST BOARD (SMS NOT SEND TO THE FARMER): 1. CHECK THE GSM/SIM MODULE CONNECTIVITY (INSERTED SIM DIRECTION) AND SIM VALIDITY -> CORRECT CONNECTIVITY / EXTEND VALIDITY. 2. CHECK THE FARMER DETAILS ENTERED ARE CORRECT -> PROPERLY ENTER DETAILS. 3. CHECK THE ENTERED SMS ENABLE SETTINGS ARE CORRECT -> ENTER CORRECT SETTINGS.
ecod_dpst_cloud_error | Tested result to cloud updation error shown Cannot upload to cloud | To fix ECOD-DPST BOARD (TESTED RESULT TO CLOUD UPDATION ERROR): 1. CHECK THE WIFI/GSM CONNECTIVITY -> ENTER PROPER SETTINGS. 2. CHECK THE ENTERED CLOUD DETAILS ARE CORRECT -> ENTER PROPER CLOUD ADDRESS. 3. CHECK THE SAME MACHINE ID AND MODEL DETAILS ARE ENTERED IN CLOUD APPLICATION -> ENTER IN WEB APPLICATION.
ecod_dpst_sd_card_save_error | Chart file not shown Farmer details not saved Collection not saved | To fix ECOD-DPST BOARD (CHART/FARMER/COLLECTION NOT SAVED): 1. CHECK THE SD CARD FILE FORMAT TYPE- (USE FAT SYSTEM) -> FORMAT THE SD CARD AT FAT FILE SYSTEM USE THE COMPUTER. 2. CHECK THE SD CARD IS DETECTED THE MACHINE (READ/WRITABLE) -> FORMAT THE SD CARD USE THAT MACHINE.
ecod_dpst_sample_test_timeout | Sample testing time the count continuously counted above 50 seconds Test taking too long | To fix ECOD-DPST BOARD (SAMPLE TESTING TIME CONTINUOUSLY COUNTED ABOVE 50 SEC): 1. CHECK THE DPST BOARD TO ANALYZER MAIN BOARD INTERFACE CABLE -> REPLACE CABLE. 2. CHECK THE DPST BOARD FUNCTIONALITY -> REPLACE DPST BOARD. 3. CHECK THE ANALYZER MAIN BOARD FUNCTIONALITY -> REPLACE MAIN BOARD.
ecod_dpst_communication_failure | DPST board to analyzer main board communication not properly working | To fix ECOD-DPST BOARD (COMMUNICATION NOT PROPERLY WORKING): 1. CHECK THE ANALYZER MAINBOARD MACHINE MODEL SELECTION IS 'ECOD'.
ecod_dpst_pendrive_failure | Pendrive detection failure USB drive not reading | To fix ECOD-DPST BOARD (PENDRIVE DETECTION FAILURE): 1. CHECK THE PENDRIVE FUNCTIONALITY AND FILE FORMAT IS FAT FILE SYSTEM -> CHANGE FILE FORMAT TO FAT. 2. CHECK THE BACKPANEL BOARD FUCTIONALITY -> REPLACE BACKPANEL BOARD.
ecod_dpst_external_keyboard_failure | External keyboard not detected | To fix ECOD-DPST BOARD (EXTERNAL KEYBOARD NOT DETECTED): 1. CHECK THE DPST BOARD TO BACKPANEL BOARD FRC CABLE QUALITY AND CONNECTOR DIRECTION -> REPLACE OR RECONNECT FRC CABLE. 2. CHECK THE BACKPANEL BOARD FUCTIONALITY -> REPLACE BACKPANEL BOARD. 3. CHECK THE DPST BOARD FUNCTIONALITY -> REPLACE DPST BOARD.
ecod_dpst_external_display_failure | External display not working | To fix ECOD-DPST BOARD (EXTERNAL DISPLAY NOT WORKING): 1. CHECK THE DPST BOARD TO BACKPANEL BOARD FRC CABLE QUALITY AND CONNECTOR DIRECTION -> REPLACE OR RECONNECT FRC CABLE. 2. CHECK THE BACKPANEL BOARD FUCTIONALITY -> REPLACE BACKPANEL BOARD. 3. CHECK THE DPST BOARD FUNCTIONALITY -> REPLACE DPST BOARD. 4. CHECK THE EXTERNAL DISPLAY INTERFACING CABLE -> REPLACE EXTERNAL DISPLAY INTERFACING CABLE.
ecod_dpst_printer_failure | Printer not working | To fix ECOD-DPST BOARD (PRINTER NOT WORKING): 1. CHECK THE PRINTER POWER SUPPLY - 8 TO 9V -> CHANGE THE POWER CABLE OR PRINTER SUPPLY BOARD. 2. CHECK THE PRINTER FUNCTIONALLY WORKING - SELF TEST PRINTOUT -> CHANGE THE PRINTER. 3. CHECK THE PRINTER SETTINGS -> CHANGE THE SETTINGS. 4. CHECK THE COMMUNICATION CABLE -> CHANGE THE CABLE.
pump_not_working | Pump not working Motor completely dead | To fix PUMP (NOT WORKING): 1. CHECK THE ANALYZER MAINBOARD BREATHING LIGHT IS WORKING -> REPLACE THE ANALYZER BOARD. 2. CHECK THE MOTOR CABLE AND CONNECTOR CONNECTIVITY -> REPLACE THE MOTOR. 3. CHECK THE MOTOR MOTION SENSING BOARD CABLE AND CONNECTOR -> REPLACE THE CABLE. 4. CHECK THE MOTOR MOTION SENSOR BOARD FUNCTANALITY -> REPLACE THE MOTOR MOTION SENSOR BOARD. 5. CHECK THE MOTOR DRIVER MODULE QUALITY -> REPLACE THE MOTOR DRIVER MODULE. 6. CHECK THE ANALYZER MAINBOARD FUNCTIONALITY -> REPLACE THE ANALYZER MAINBOARD.
pump_nut_bolt_clamp_hit | Pump nut and bolt move to oneside only and hit into the clamp | To fix PUMP (NUT AND BOLT MOVE TO ONESIDE AND HIT CLAMP): 1. CHECK THE MOTOR MOTION SENSING BOARD CABLE AND CONNECTOR -> REPLACE THE MOTOR MOTION SENSING BOARD CABLE. 2. CHECK THE MOTOR MOTION SENSING BOARD FUNCTIONALITY -> REPLACE THE SENSING BOARD. 3. CHECK THE ANALYZER MAINBOARD FUNCTIONALITY -> REPLACE THE ANALYZER MAINBOARD.
pump_sense_error | Pump sense error | To fix PUMP (PUMP SENSE ERROR): 1. CHECK THE MOTOR MOTION SENSING BOARD CABLE AND CONNECTOR -> REPLACE SENSING BOARD CABLE. 2. CHECK THE MOTOR MOTION SENSING BOARD FUNCTIONALITY -> REPLACE SENSING BOARD. 3. CHECK THE ANALYZER MAINBOARD FUNCTIONALITY -> REPLACE ANALYZER MAINBOARD.
ecod_battery_low_error | Low battery error shown | To fix ECOD/ECOSV/LSES V3 BATTERY PROBLEMS (LOW BATTERY ERROR): 1. CHECK THE BATTERY VOLT ABOVE 11.5V -> REPLACE THE BATTERY. 2. CHECK THE SOLARCHARGER BOARD TO ANALYZER INTERFACE CABLE -> REPLACE THE INTERFACE CABLE. 3. CHECK THE ANALYZER MAINBOARD FUNCTIONALITY -> REPLACE THE ANALYZER MAINBOARD. 4. CHECK THE SOLAR CHARGER BOARD FUNCTANALITY -> REPLACE THE SOLAR CHARGER BOARD.
ecod_battery_full_led_error | Battery full LED indication not working | To fix ECOD/ECOSV/LSES V3 BATTERY PROBLEMS (BATTERY FULL LED NOT WORKING): 1. CHECK THE ANALYZER MAINBOARD IS COMPLETED THE WARMING PROCESS -> APPLY TO CANCELLED THE WARMING PROCESS. 2. CHECK THE VOLTAGE OF THE BATTERY ABOVE 14.3V -> CHARGE THE BATTERY AT FULL. 3. CHECK THE CHARGER BOARD TO ANALYZER MAINBOARD INTERFACING CABLE -> REPLACE CABLE. 4. CHECK VOLTAGE OF SOLAR CHARGER BOARD TO BATTERY CABLE LINE WITHOUT CONNECTING BATTERY (14.45V) -> ADJUST TRIMPOT AND SET AT 14.45V. 5. CHECK THE BATTERY FULL INDICATION LED IS WORKING -> REPLACE THE LED. 6. CHECK THE ANALYZER MAIN BOARD TO LED CABLE -> REPLACE CABLE. 7. CHECK THE ANALYZER MAINBOARD FUNCTIONALITY -> REPLACE MAINBOARD.
analyzer_mainboard_temp_set_error | T2 TEMP.SET ERROR Temperature set error | To fix ANALYZER MAINBOARD (T2/TEMP.SET ERROR): 1. CHECK THE LEAKAGE/BLOCK OF THE SAMPLE SUCKING SECTIONS. 2. CHECK THE INLET PIPE IS BROKEN OR BLOCKED -> REPLACE INLET PIPE. 3. CHECK THE INLET TUBE TO PREHEATER SILICON TUBE -> REPLACE/TIE SILICON TUBE. 4. CHECK THE PRE HEATER TO SENSOR SILICON TUBE -> REPLACE/TIE SILICON TUBE. 5. CHECK THE SENSOR TO SYRINGE SILICON TUBE -> REPLACE/TIE SILICON TUBE. 6. CHECK THE SYRINGE IS DAMAGE -> REPLACE THE SYRINGE. 7. CHECK THE BELLOW TO L-PLUG SILICON TUBE BROKEN OR BUR HOLDED -> REPLACE SILICON TUBE/BUR. 8. CHECK THE L-PLUG BROKEN AND O-RING QUALITY -> REPLACE L-PLUG/O-RING. 9. CHECK THE BELLOW QUALITY -> REPLACE BELLOW. 10. CHECK THE SENSOR INNER BOX SIDE ANY LEAKAGE/BLOCK -> REPLACE MAINBOARD AND SENSOR BOARD PAIR. 11. CHECK MAINBOARD FUNCTIONALITY -> REPLACE MAINBOARD AND SENSOR BOARD PAIR.
analyzer_mainboard_box_temp_error | BOX TEMP ERROR | To fix ANALYZER MAINBOARD (BOX TEMP ERROR): 1. CHECK THE JUMPER QUALITY OF THE MAINBOARD TO SENSOR BOARD SUPPLY LINE -> REPLACE THE JUMPER. 2. CHECK THE SENSOR BOARD FUNCTIONALITY -> REPLACE THE SENSOR BOARD. 3. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR BOARD PAIR.
analyzer_mainboard_sample_not_found | SAMPLE NOT FOUND AIR IN MILK | To fix ANALYZER MAINBOARD (SAMPLE NOT FOUND/AIR IN MILK): 1. CHECK THE SAMPLE SUCKING SECTION ANY LEAKAGE/BLOCK -> REPLACE THE DAMAGE ITEM. 2. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR BOARD PAIR.
analyzer_mainboard_plunge_in_water | PLUNGE IN WATER WATER IN SENSOR | To fix ANALYZER MAINBOARD (PLUNGE IN WATER/WATER IN SENSOR): 1. CHECK THE SAMPLE LIQUID PRESENTS IN THE SENSOR TUBE -> REMOVE THE SAMPLE LIQUID. 2. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR BOARD PAIR.
analyzer_mainboard_hot_sample_error | HOT SAMPLE ERROR Sample is too hot | To fix ANALYZER MAINBOARD (HOT SAMPLE ERROR): 1. CHECK THE TEST SAMPLE TEMPERATURE BELOW 40 DEGREE -> REPLACE THE SAMPLE. 2. CHECK THE SENSOR BOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR BOARD PAIR. 3. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR BOARD PAIR.
analyzer_mainboard_count_zero | COUNT ZERO RESULT ZERO | To fix ANALYZER MAINBOARD (COUNT ZERO / RESULT ZERO): 1. CHECK THE SENSOR BLOCK / LEAKAGE AND TEST COUNT VALUE -> CLEAN THE SENSOR. 2. CHECK THE SENSOR GAIN SET AT PROPER VALUE -> SET THE NORMAL GAIN SETTINGS. 3. CHECK THE SENSOR BOARD FUNCTIONALITY -> REPLACE THE SENSOR BOARD. 4. CHECK MAIN BOARD FUNCTIONALITY -> REPLACE MAINBOARD AND SENSOR PAIR.
analyzer_mainboard_reading_variation | READING VARIATION Inconsistent readings | To fix ANALYZER MAINBOARD (READING VARIATION): 1. CHECK THE SENSOR AND SILICONE TUBES ARE CLEANED -> CLEAN THE SENSOR AND SILICONE TUBES. 2. CHECK THE SENSOR GAIN SET AT PROPER VALUE -> SET THE NORMAL GAIN SETTINGS. 3. CHECK THE PROPER OFFSET APPLY AND TEST READING -> REPLACE MAINBOARD AND SENSOR BOARD PAIR. 4. CHECK MAIN BOARD FUNCTIONALITY -> REPLACE MAINBOARD AND SENSOR PAIR.
analyzer_mainboard_fat_in_water | FAT SHOWN IN WATER | To fix ANALYZER MAINBOARD (FAT SHOWN IN WATER): 1. CHECK THE ANALYZER OTHER MILK SAMPLES TEST READINGS ARE 'OK' -> IF NOT OK CHECK THE CALIBRATION DETAILS. 2. CHECK THE SENSOR AND TUBE WELL CLEANED -> TO CLEAN THE SENSOR AND TUBE, APPLY WATER ZERO OPERATION AND TEST AGAIN. 3. ENSURE THE PAIRED MAIN BOARD IS FUNCTIONALLY OK OTHERWISE BOARD REPLACE.
analyzer_mainboard_lcd_computer_mode | LCD CONTINIOUSLY SHOWING COMPUTER MODE | To fix ANALYZER MAINBOARD (LCD CONTINIOUSLY SHOWING COMPUTER MODE): 1. CHECK THE FRONTPANEL KEYPAD KEYS QUALITY -> REPLACE THE KEYPAD. 2. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR BOARD PAIR.
analyzer_mainboard_lcd_version | LCD CONTINIOUSLY SHOWING VERSION | To fix ANALYZER MAINBOARD (LCD CONTINIOUSLY SHOWING VERSION): 1. CHECK THE FRONTPANEL KEYPAD KEYS QUALITY -> REPLACE THE KEYPAD. 2. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR BOARD PAIR.
analyzer_mainboard_fuse_burn | WARMING STARTING TIME MACHINE WILL OFF FUSE BURN | To fix ANALYZER MAINBOARD (WARMING STARTING TIME MACHINE WILL OFF/FUSE BURN): 1. CHECK THE MAINBOARD FUNCTIONALITY - MOSFET FAILURE -> REPLACE MAIN BOARD AND SENSOR BOARD PAIR. 2. CHECK THE SENSOR BOARD COIL IS BURN -> REPLACE MAIN BOARD AND SENSOR BOARD PAIR. 3. CHECK THE ON/OFF SWITCH FUNCTIONALITY AND FUSE CARRER & FUSE IS TIGHT -> REPLACE SWITCH OR RE-SET FUSE AND CARRIER. 4. CHECK THE POWER CABLE OR CONNECTOR IS LOOSE -> REPLACE FAULTY ITEM.
analyzer_mainboard_pendrive_keyboard_undetected | PENDRIVE AND KEYBOARD NOT DETECTED | To fix ANALYZER MAINBOARD (PENDRIVE AND KEYBOARD NOT DETECTED): 1. CHECK THE MAINBOARD USB SETTING IS CORRECT -> CORRECT THE USB SETTINGS. 2. CHECK THE USB CONNECTOR AND INTERFACING CABLE -> REPLACE THE CONNECTOR OR CABLE. 3. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR BOARD PAIR.
analyzer_mainboard_wifi_gsm_error | WIFI/GSM MODULE ERROR SHOWN | To fix ANALYZER MAINBOARD (WIFI/GSM MODULE ERROR SHOWN): 1. CHECK THE WIFI/GSM SETTINGS DETAILS -> ENTER REQUIRED SETTINGS. 2. CHECK THE MAINBOARD TO MODULES OR BACKPANEL FRC CABLE -> REPLACE OR RECONNECT AT CORRECT DIRECTION. 3. CHECK THE WIFI/GSM MODULE BREATHING LED IS WORKING -> REPLACE MODULE. 4. CHECK THE ANTENNA SMA CABLE CONNECTION AND CONNECTOR -> REPLACE ANTENNA AND SMA CABLE.
analyzer_mainboard_frontpanel_keypad | FRONTPANEL KEYPAD NOT WORKING | To fix ANALYZER MAINBOARD (FRONTPANEL KEYPAD NOT WORKING): 1. CHECK THE KEYPAD, INTERFACING CABLE AND CONNECTOR QUALITY -> REPLACE THE KEYPAD, INTERFACING CABLE AND CONNECTOR. 2. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR PAIR.
analyzer_mainboard_lcd_blank | LCD NOT WORKING BLANK DISPLAY | To fix ANALYZER MAINBOARD (LCD NOT WORKING - BLANK DISPLAY): 1. CHECK THE LCD, INTERFACE CABLE AND CONNECTOR QUALITY -> REPLACE THE LCD, INTERFACING CABLE AND CONNECTOR. 2. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR PAIR.
analyzer_mainboard_rtc_not_running | RTC TIME NOT RUNNING | To fix ANALYZER MAINBOARD (RTC TIME NOT RUNNING): 1. CHECK THE RTC MODULE IC CRYSTAL AND FILTER CAPACITORS -> REPLACE CRYSTAL AND CAPACITORS. 2. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR PAIR.
analyzer_mainboard_computer_output_not_present | COMPUTER OUTPUT NOT PRESENT | To fix ANALYZER MAINBOARD (COMPUTER OUTPUT NOT PRESENT): 1. CHECK THE DATAOUTPUT FORMAT OF THE ANALYZER SETTINGS -> SET REQUIRED SETTINGS VALUES. 2. CHECK THE MACHINE COMPUTER OUT TO COMPUTER CABLE CONNECTIVITY -> REPLACE CABLE. 3. CHECK THE MAINBOARD FUNCTIONALITY -> REPLACE THE MAINBOARD AND SENSOR PAIR.
analyzer_mainboard_rtc_time_not_saved | RTC TIME NOT SAVED RESTART TIME TIME CHANGED | To fix ANALYZER MAINBOARD (RTC TIME NOT SAVED): 1. CHECK THE CMOS BATTERY VOLTAGE ABOVE 2V -> REPLACE THE CMOS BATTERY. 2. CHECK THE RTC MODULE AND MAINBOARD FUNCTIONALITY -> REPLACE THE RTC MODULE OR MAINBOARD PAIR.
analyzer_mainboard_sms_not_send | SMS NOT SEND TO THE FARMER | To fix ANALYZER MAINBOARD (SMS NOT SEND TO THE FARMER): 1. CHECK THE GSM/SIM MODULE CONNECTIVITY AND SIM VALIDITY -> CORRECT CONNECTIVITY AND EXTEND VALIDITY. 2. CHECK THE FARMER DETAILS ENTERED ARE CORRECT -> PROPERLY ENTER FARMER DETAILS. 3. CHECK THE ENTERED SMS ENABLE SETTINGS ARE CORRECT -> ENTER CORRECT SETTINGS.
analyzer_mainboard_cloud_error | TESTED RESULT TO CLOUD UPDATION ERROR SHOWN | To fix ANALYZER MAINBOARD (TESTED RESULT TO CLOUD UPDATION ERROR): 1. CHECK THE WIFI/GSM CONNECTIVITY -> ENTER PROPER SETTINGS. 2. CHECK THE ENTERED THE CLOUD DETAILS ARE CORRECT -> ENTER PROPER CLOUD ADDRESS AND DETAILS. 3. CHECK THE SAME MACHINE ID AND MODEL DETAILS ARE ENTERED IN THE CLOUD APPLICATION -> ENTER ID AND MODEL IN WEB APP.
analyzer_mainboard_rate_not_taken | RATE NOT TAKEN FROM THE CHART NOT VIEW | To fix ANALYZER MAINBOARD (RATE NOT TAKEN FROM THE CHART / NOT VIEW): 1. CHECK THE RATE CHART SETTINGS ENTERED PROPERLY -> ENTER THE CORRECT SETTINGS.
analyzer_mainboard_farmer_details_not_shown | FARMER DETAILS NOT SHOWN | To fix ANALYZER MAINBOARD (FARMER DETAILS NOT SHOWN): 1. CHECK THE FARMER DETAILS ARE ENTERED CORRECTLY -> ENTER THE CORRECT DETAILS.
analyzer_mainboard_weighing_scale | WEIGHING SCALE NOT WORKING | To fix ANALYZER MAINBOARD (WEIGHING SCALE NOT WORKING): 1. CHECK THE REQUIRED SETTINGS IS APPLIED -> SET REQUIRED SETTINGS VALUES. 2. CHECK THE MAIN BOARD TO BACKPANEL BOARD FRC CABLE -> REPLACE CABLE AND RECONNECT. 3. CHECK THE BACKPANEL PCB FUNCTANALITY -> REPLACE BACKPANEL PCB. 4. CHECK THE MAIN BOARD FUNCTANALITY -> REPLACE MAINBOARD AND SENSOR PAIR. 5. CHECK THE WEIGHING SCALE INTERFACING CABLE -> REPLACE CABLE.
analyzer_mainboard_external_display | EXTERNAL DISPLAY NOT WORKING | To fix ANALYZER MAINBOARD (EXTERNAL DISPLAY NOT WORKING): 1. CHECK THE MAIN BOARD TO BACKPANEL BOARD FRC CABLE -> REPLACE CABLE AND RECONNECT. 2. CHECK THE BACKPANEL BOARD FUCTIONALITY -> REPLACE BACKPANEL BOARD. 3. CHECK THE MAIN BOARD FUNCTIONALITY -> REPLACE MAIN BOARD. 4. CHECK THE EXTERNAL DISPLAY INTERFACING CABLE -> REPLACE CABLE.
others_analyzer_not_working | ANALYZER NOT WORKING Analyzer is completely dead | To fix OTHERS (ANALYZER NOT WORKING): 1. CHECK THE FUSE -> REPLACE THE FUSE. 2. CHECK THE ADAPTER / BATTERY VOLTAGE ABOVE 11.5V -> REPLACE THE ADAPTER / BATTERY. 3. CHECK THE ON/OFF SWITCH FUNCTIONALITY -> REPLACE THE ON/OFF SWITCH. 4. CHECK THE INTERNAL BARDS SUPPLY LINE IS SHORTCIRCUIT -> REPLACE THE SHORTED BOARD. 5. CHECK THE INTERNAL SUPPLY LINE OPEN/LOOSE FAULT -> REPLACE THE FAULTY CONNECTOR/CABLES.